## Sheet1
For Goods Offered From Within the Philippines
Contract ID. No.: |  |  | 24GDH0007
Name of Contract: |  |  | Purchase and Delivery of Laboratory Testing Equipment/Apparatus and Supplies of Quality Assurance Section-DPWH Laguna 1st District Engineering Office, Sta. Cruz, Laguna
Location : |  |  | DPWH Laguna 1ˢᵗ District Engineering Office, Sta. Cruz, Laguna
1 | 2 |  |  | 3 | 4 |  | 5 | 6 | 7 | 8 | 9 | 10
Item | Description |  |  | Country of Origin | Quantity |  | Unit Price EXW per item | Transportation and all other costs incidental to delivery, per item | Sales and other taxes payable if Contract if awarded, per item | Cost of incidental Services,If applicable, per item | Total Price, per unit | Total Price delivered Final Destination
[11 empty] | (col.5+6+7+8) | (col 9)x (col 4)
1 |  | Mortar Mixer |  |  | 1 | unit
Capacity: 5qt (4.7L)
Operating Speed: 139,285, or 591 rpm (selectable
Electical:230V, 50Hz, 1.4 Amps
Dimensions (WxDxH): 10.5x15x17in (267x381x432mm)
100-120/60/1 Mixer with Bowl & Stainless Steel; includes
enhanced bowl support System; ASTM concrete lab
configuration mixer-ASTM Approved,Planetary, Bench,
gear driven, transmission, #10 paper attachment hub,
manual bowl lift, enamel housing, stainless steel bowl,
Stainless steel "B" beater, 1/6 hp, cord with plug
*Selectable operating speeds so you can mix at a rate
that suits your application
*Compliant with ASTM C305, C227, and C109 (AASHTO
162 and T 106) and other tests for mortar and cement
*Locking hand lever easily raises and lowers the bowl
CBR Accessories
2 |  | CBR Mold 6" x 7" |  |  | 3 | pc
Molds-The molds shall be cylindrical in shape, made of
metal, with an internal diameter of 152.40 ± 0.66 mm
(6.0± 0.026 in.) and a height of 177.80 ± 0.46 mm (7.0
± 0.018 in), and provided with an extension collar
approximately 50 mm (2.0 in) in height and a perforated
base plate that can be fitted to either end of the mold.
Sub-Total P
|  | Amount Forwarded | [8 empty] | P
3 |  | Compaction Hammer 10 lbs |  |  | 1 | pc
Manually Operated-Metal rammer with a mass of 4.536
± 0.009 kg (10.00 ± 0.02 lb) (Note 3) and having a flat
circular face of 50.80 mm (2.000 in.) diameter with a
manufacturing tolerance of ±0.25 mm (±0.01 in.). The in-
service diameter of the flat circular face shall be not less
than 50.42 mm (1.985 in.). The rammer shall be
equipped with a suitable guide sleeve to control the
height of drop to a free fall 457 ± 2 mm (18.00 ± 0.06 in.)
above the elevation of the soil. The guide sleeve shall
have at least four vent holes, no smaller than 9.5-mm
(0.375 in.) diameter, spaced approximately 90 degrees
(1.57 rad) apart from each end, and shall provide
sufficient clearance so the free fall of the rammer shaft
and head is unrestricted
4 |  | Spacer Disc 5-7/8 |  |  | 3 | pc
Spacer Disk-A circular spacer disk made of metal 150.8
± 0.8 mm (515/16 ± 1/32 in.) in diameter and 61.37 ±
0.25 mm (2.416 ± 0.01 in) in height.
Note 1 -When using molds having a height of 177.80
mm (7.0 in.), a spacer disk height of 61.37 mm
(2.416 in.) is needed to obtain a thickness of
compacted specimen that conforms to the
thickness: 116.43 mm (4.584 in.) of specimen
in T 99 and T 180
Sub-Total P
|  | Amount Forwarded | [8 empty] | P
5 |  | Slotter Surcharge Weights |  |  | 3 | pc
Slotter metal weights 149.2±1.6 mm (57/8±1/6 in.) in
diameter and 2.27 ± 0.04kg. (5±0.10 lb) weight
6 |  | Annual Surcharge Weights |  |  | 3 | pc
Annular metal weight with a center hole approximately
54.0 mm (21/8 in.) in diameter, 2.27± 0.04kg (5 ± 0.10 lb)
weight
7 |  | Straight Edge 12" Long (Stainless) |  |  | 1 | pc
8 |  | Swell Plate (BRASS) |  |  | 3 | pc
Swell plate with adjustable stem and a tripod support for
a dial indicator. The swell plate is made of metal, 149.2
± 1.6 mm )57/8 ± 1/16 in.) in diameter and is perforated
with 1.6 mm(1/16in.) diameter holes. The tripod used
to support the dial indicator is arranged to fit the mold
extension collar.
9 |  | Analog Dial Indicator, 0.01mm-0.10mm |  |  | 3 | pc
10 |  | Fumehood |  |  | 1 | unit
With complete scrubber Stainless (304) and duct system
(Fiber Glass) including installation for Fuming Acid
digestion(HCI, H2S04, HN03, and other acids of lower
strength)and Ammonia Vapour. Overhead LED light
Strip or fluorescent lamp attached to inside top of hood
with corrossion and moisture protection, 1000 lux (100 foot
candles)
Sub-Total P
|  | Amount Forwarded | [8 empty] | P
Minimum internal fume hood/working area dimension =
(width x depth x height) 1397mm x 736mm x 1499mm
Maximum height should not exceed laboratory ceilling)
Worktop = Made of Granite, Heat, Flame and corrosion
resistant (such as epoxy resin homogenous solid) with spill
drainage of spill stopper; Volume of air treated: min = 690 m3/hr
Air face velocity: variable up to at least 0.5m/s t full
sash opening. Voltage: 220/240/220-240V);
Laminated tempered glass type sash or materials of
similar or better performance, with counterweight or lock
to allow sash to stay at preferred height, corrosion
resistent; Electric socket outside near the hood opening,
minimum of 2 outlet. Controller for air velocity including
calibration upon installations; Alarm/alert system for filter
or scrubber saturation. Air velocity, fan failure and sash
height for ductless and ducted fume hood. Chemical
resistant (epoxy resin coated or higher performance)
cabinet at the bottom of the fume hood; at least
ASHRAE 110 certified and /or EN14175 certified;
Delivery, installation and commissioning of ductless fume
Hood or ducted fume hood with complete scrubber and
duct system; Exclusive distributorship or authorized
Distributor with local after sales support and parts
availability for at least 5 years form purchase of fume
hood; Preventive maintenance of fume hood after 6
mos. Inclusive of scrubber maintenance.
11 |  | Digital Caliper 8" |  |  | 2 | pc
12 |  | Glass Pipette 10ml with aspirator |  |  | 1 | pc
Sub-Total P
|  | Amount Forwarded | [8 empty] | P
13 |  | Graduated Cylinder 10 ml. |  |  | 1 | pc
14 |  | Sieve Brass - 3" (75mm) x 8" (200mm) dia. |  |  | 1 | pc
15 |  | Sieve Brass - #80 (75mm) x 8" (200mm) dia. |  |  | 1 | pc
16 |  | CBR Soaking Tank |  |  | 1 | unit
- Galvanized Steel (1.5mm thickness)
- Size: 55 cm x 33cm x 35cm
17 |  | Working Table |  |  | 1 | unit
- 2x2 (1.5mm thickness) Square Tubing Galvanized
- granite (radiant white quartz)
- Size: 3ft x 9ft. X 82cm height
- Heavy Duty
18 |  | Working Table |  |  | 3 | unit
- 2x2 (1.5mm thickness) Square Tubing Galvanized
- granite (radiant white quartz)
- Size: 3ft x 32inches X 75cm height
- Heavy Duty
19 |  | Laboratory Chemical Cabinet |  |  | 1 | unit
- Galvanized Steel with glass door
Height: 1.5m
Width: 0.80m
Depth: 0.40m
20 |  | Waterproof Digital Walking Measure |  |  | 2 | unit
- 100km Type, Minimum Display - 0.05m
21 |  | Meter Tape, 8m/26" |  |  | 2 | unit
Sub-Total P
|  | Amount Forwarded | [8 empty] | P
22 |  | Heavy Duty Screw driver set, Pliers, Locking |  |  | 1 | set
Pliers, Wrench, Hammer, Mechanical Tool Set &
Tool Box
x-x-x-x-x-x-x-x-x-x-x-x-x-x-x-x-x
| Bid Validity: |  | 120 calendar days |  | Total Amount of Bid:
P
[5 empty] | (in words) | [6 empty] | (in figure)
Name :
Legal Capacity:
Signature:
Duly Authorized to sign bid for and on behalf of:
Date: